OX800 metrix

_____________________________________________________________________________

Introduction:

This  instrument  module  provides RS232  programming  support

for  the ox800 metrix scope (OX800,OX801,OX830 and OX830D) .  It contains

functions for opening and configuring the scope.

_____________________________________________________________________________

Assumptions:

To successfully use this module, the following conditions must be met:

- the ox800 is connected to the RS232 (9600,8,N,1 with XON/XOFF protocole)

25 SUB D                             9 SUB D

PC           OX800                PC              OX800

FG   1                         *  DCD   1 <-\

TXD  2 ----------> 3 RXD          RXD   2 <-I----------- 2 TXD

RXD  3 <---------- 2 TXD          TXD   3 --I----------> 3 RXD

*  RTS  4 --\                     *  DTR   4 --I

*  CTS  5 <-I                        SG    5 --I---------- 7 SG

*  DSR  6 <-\                     *  DSR   6 <-I

SG   7 --I--------- 7 SG       *  RTS   7 --\

*  DCD  8 <-I                     *  CTS   8 <-I

*  DTR 20 <-I                        FG    9

* : (4,5 and 20,8,6 tied together)   * : (7,8 and 4,1,6 tied together)

_____________________________________________________________________________

LabWindows Libraries:

In  order to use this instrument module outside of the LabWindows envi-

ronment, it must be linked with the following LabWindows libraries:

- Formatting and I/O

- RS-232

_____________________________________________________________________________

Error Codes:

All  error codes are  returned  in a global integer variable whose name

depends on the language being used:

For QuickBASIC : ox800.err%

For C          : ox800_err

A program should examine this variable after each function call is made

to determine if any errors occurred.

_____________________________________________________________________________

Function Tree Layout:

OX800 metrix

Initialize

Instrument setup

Configure

Display settings

Chanel  settings

TimeBase

Trigger

Close

_____________________________________________________________________________

Function           : Initialize

QuickBASIC syntax  : ox800.init(rs_port%)

C syntax           : void ox800_init(rs_port)

int rs_port;

Purpose            :  Open and configure the RS232 port , configure the

instrument with default values .

Identifies the instrument (OX800_51 or OX800_52);

If instrument OX800_51 then a Delay (0.01 sec) is

included before sendings commands to the instrument.

The RS232 default configuration is :

9600 bauds, no parity, 8 bits word, 1 stop bit

with XON/XOFF protocol.

A 1.0 Seconds timeout is initialized.

Parameter List     :

rs_port :   Serial port number (0 -> 8)

QuickBASIC example :  CALL ox800.init(1)

C example          :  ox800.init(1)

Error Codes        :  The valid error codes for this function are :

0   :       success

-1   :       rs_port value out of range

Others  :       see rs232_err (RS232 LabWindows Library)

_____________________________________________________________________________

Function           : Remote (for OX800_52 only)

QuickBASIC syntax  : ox800.set.remote(param%)

C syntax           : void ox800_set_remote(param)

int param;

Purpose            :  Sets the instrument in local or remote control .

Parameter List     :

param :   0 : local  (front pannel enabled)

1 : remote (front pannel disabled)

QuickBASIC example :  CALL ox800.set.remote(1)

C example          :  ox800.set.remote(1)

Error Codes        :  The valid error codes for this function are :

0   :       success

-2   :       param value out of range

-3   :       Remote function not available with this instrument

Others  :       see rs232_err (RS232 LabWindows Library)

_____________________________________________________________________________

Function           :  Instrument Setup

QuickBASIC syntax  :  ox800.setup(config%,file$)

C syntax           :  void ox800_setup(config,file)

int config;

char *file;

Purpose            :  Read/write instrument setup

from/to file

Parameter List     :

config  :   0   :   Save

1   :   Recall

File    :   file name

QuickBASIC example :  CALL ox800.setup(1,"setup")

C example          :  ox800_setup(1,"setup")

Error Codes        :  The valid error codes for this function are :

0  :  Success

-1  :  Setup value out of range

-2  :  File doesn't exist

230  :  Write Error

231  :  Read Error

232  :  Instrument not initialized

Others :  See rs232err values

_____________________________________________________________________________

Function           :  Display settings

QuickBASIC syntax  :  CALL ox800.display(mod%)

C syntax           :  void ox800_display(mod)

int mod;

Purpose            :  Configure the Vertical mode and control

the Beam Find .

Parameter List     :

mode: Vertical display mode

0 : CH1

1 : ALT         ( default value )

2 : CHOP

3 : ADD

4 : CH2

5 : XY

6 : TEST        (OX800 and OX830 only)

QuickBASIC example :  CALL ox800.display(0)

C example          :  ox800_display(0);

Error Codes        :  The valid error codes for this function are :

0  :       Success

-1  :       Vertical Mode out of range

-2  :       Beam Finder out of range

230  :       RS232 write error

232  :       Instrument not initialized

_____________________________________________________________________________

Function           :  Beam finder

QuickBASIC syntax  :  ox800.beam.finder(beam%)

C syntax           :  ox800_beam_finder(beam)

int beam;

Purpose            :  Sets beam finder

Parameter List     :

beam :

0 : Off

1 : On

QuickBASIC example :  CALL ox800.beam.finder(1)

C example          :  ox800_beam_finder(1);

Error Codes        :  The valid error codes for this function are :

0  :       Success

-1  :       Vertical Mode out of range

-2  :       Beam Finder out of range

230  :       RS232 write error

232  :       Instrument not initialized

_____________________________________________________________________________

Function           :  Channel settings

QuickBASIC syntax  :  CALL ox800.channel(channel_num%,coupling%,sensibility%)

C syntax           :  void ox800_channel(channel_num,coupling,sensibility)

int channel_num;

int coupling;

int sensibility;

Purpose            :   Sets the Channel , its coupling and volts/division .

Parameter List     :  channel_num :    channel number

0  : CH1      (default value)

1  : CH2

coupling   :

0  :    AC       (default value)

1  :    DC

2  :    GND

Sensibility :

13  :     1  mV  (not available on OX801)

12  :     2  mV

11  :     5  mV

10  :    10  mV

9  :    20  mV

8  :    50  mV

7  :     0.1 V    (default value)

6  :     0.2 V

5  :     0.5 V

4  :     1   V

3  :     2   V

2  :     5   V

1  :    10   V

0  :    20   V

Note : Range x 10 for OX830D (from 10mV to 200V/Div)

QuickBASIC example :  CALL ox800.channel(1,0,0)

C example          :  ox800_channel(1,0,0);

Error Codes        :  The valid error codes for this function are :

0  :  Success

-1  :  Channel value out of range

-2  :  Coupling value out of range

-3  :  Sensibility value out of range

-4  :  Invert CH2 value out of range

-5  :  Not allowed with Vertical mode TEST

(See Display settings)

230  :  Write Error

232  :  Instrument not initialized

_____________________________________________________________________________

Function           :  Invert CH2

QuickBASIC syntax  :  CALL ox800.chan.invert(inversion%)

C syntax           :  void ox800_chan_invert(inversion)

int inversion;

Purpose            :  Sets Invert CH2 .

Parameter List     :

inversion :

0  :  Off       (default value)

1  :  On

QuickBASIC example : CALL ox800.chan.invert(0)

C example          : ox800_chan_invert(0)

Error Codes        :  The valid error codes for this function are :

0  :  Success

-1  :  Channel value out of range

-2  :  Coupling value out of range

-3  :  Sensibility value out of range

-4  :  Invert CH2 value out of range

-5  :  Not allowed with Vertical mode TEST

(See Display settings)

230  :  Write Error

232  :  Instrument not initialized

_____________________________________________________________________________

Function           :  Time base

QuickBASIC syntax  :  ox800.timebase(expend%,range%,delay%)

C syntax           :  void ox800_timebase(expend,range,delay)

int expend;

int range;

int delay;

Purpose            :  Sets the Time base parameters of the

oscilloscope .

Parameter List     :

expend :  x10

0 :     Off     (default value)

1 :     On

range :   From 0.5�S to 200 mS

0   :   200   mS

1   :   100   mS

2   :    50   mS

3   :    20   mS

4   :    10   mS

5   :     5   mS

6   :     2   mS

7   :     1   mS

8   :     0.5 mS

9   :     0.2 mS

10   :     0.1 mS   (default value)

11   :    50   �S

12   :    20   �S

13   :    10   �S

14   :     5   �S

15   :     2   �S

16   :     1   �S

17   :     0.5 �S

delay : state of the delayed time base :

0   : Off     (default value)

1   : Search

2   : Delay

QuickBASIC example :  CALL ox800.timebase(0,7,0)

C example          :  ox800_timebase(0,7,0);

Error Codes        :  The valid error codes for this function are :

0   :      Success

-1   :      x10 value out of range

-2   :      Range value out of range

-3   :      Delay value out of range

-4   :      Not allowed with Vertical mode TEST

(See Display settings)

-5   :      Not allowed with Vertical mode XY

(See Display settings)

230   :      RS232 write error

232   :      Instrument not initialized

_____________________________________________________________________________

Function           :  Trigger

QuickBASIC syntax  :  CALL ox800.trigger(ptp%,slope%,mod_auto%,source%,coupling%)

C syntax           :  void ox800_trigger(ptp,slope,mod_auto,source,coupling)

int ptp;

int slope;

int mod_auto;

int source;

int coupling;

Purpose            :  Sets the Trigger parameters as Peak to Peak mode , Slope ,

Auto trigger mode , source and coupling .

Parameter List     :

ptp : Peak to Peak mode

0 : Off

1 : On      (default value)

slope :

0 : Positive  (default value)

1 : Negative

mod_auto : Auto trigger mode

0 : Off

1 : On      (default value)

source :

0 : CH1     (default value)

1 : CH2

2 : ALT

3 : LINE

4 : EXT

Coupling :

0 : DC       (default value)

1 : AC

2 : LF

3 : HF

4 : TVH

5 : TVV

QuickBASIC example :  CALL  ox800.trigger(0,0,0,0,0)

C example          :  ox800_trigger(0,0,0,0,0);

Error Codes        :  The valid error codes for this function are :

0  :     Success

-1  :     Ptp value out of range

-2  :     Slope value out of range

-3  :     Auto value out of range

-4  :     Source value out of range

-5  :     Coupling value out of range

-6  :     Not allowed with Vertical mode TEST

(See Display settings)

-7  :     Not allowed with Vertical mode XY

(See Display settings)

230  :     RS232 write error

232  :     Instrument not initialized

_____________________________________________________________________________

Function           :  close

QuickBASIC syntax  :  CALL ox800.close()

C syntax           :  void ox800_close()

Purpose            :  Close the instrument

(The instrument must be reinitialized to use

it again).

Parameter List     :  NONE

QuickBASIC example :  CALL ox800.close()

C example          :  ox800_close();

Error Codes        :  The valid error codes for this function are :

0 :       Success

221 :       Unable to close instrument

232 :       Instrument not initialized

_____________________________________________________________________________